## Comparison
Name | WD-WUR v5 | MooLogger | Guardian Gascard | GT5000 | ASCH401 | Gasera one pulse | Simple scan | CH4/CO2/H2O analyser | Agscent Air
Company | Carltech | Tecnosens | Edinburgh sensors | Gasmet | Arcoflex | Gasera | C-lock | Nikira labs | Agscent
Country of origin | The Netherlands | Italy | Great-Britain | Finland | Australia | Finland | US | USA | AUS/USA
Range CO2 (ppm) | 0-50.000 | 0-50.000 | 0-50.000 | Adjustable | Included in new version | 0,3-30.000 | 0-20.000 | 100-20.000 | 0-20.000
Range CH4 (ppm) | 0-2.000 | 0-5.000 | 0-10.000 | adjustable | 50-5.000 | 1-100.000 | 0-40.000 | 1-20.000 | 0-40.000
02 sensor | Optional | Optional | No | Optional | No | Optional | Optional | No | No
H2 sensor | No | Optional | No | No | No | Optional | Optional | No | No
H20 sensor | No | Yes | No | Optional | No | Yes | No | Optional | No
Sensor type | NDIR dual beam; dual to quad wavelengths | NDIR for CH4 and CO2 Electrochemical cells for other optional gasses | CH4 NDIR; CO2 pseudo dual beam NDIR | FTIR | MQ-4 heated elektrode sensor | Photoacoustic detection, broadband IR source and narrow band optical filters | NDIR for CO2 and TDL for CH4 | Cavity Enhanced Absorption Spectroscopy | TDL for CH4 and NDIR for CO2
Response time (sec) | 1.5 – 15 | 2 | 5 | 30 | Almost instant | 15 - 60 |  | <1Hz, 1/e | <1.8 for CH4 and <30s for CO2
Data storage time (s) | 2025-05-01 | 1 | 1-5 | 1 | Almost instant | Settable | 1 | 1Hz | 1
Sample pump flow/gas flow | 0.1 – 1.5 l/min | 0,9-1L/min | Depends on build, at least 1L/min | 1.5L/min | Not applicable | Approx 0.6 L/min | 1L/min | 350 sccm | 1L/min, settable
Repeatability | <±0,3% of reading | < 0,25% F.S for CH4; < 0,4% F.S for CO2; | Not defined under 10°c |  |  | <1% |  | Precision: CH4: < 2 ppb (1σ, 1s) / < 1 ppb (1σ, 10s); CO2: < 0.5 ppm (1σ, 1s) / < 0.2ppm (1σ, 10s); H2O: < 100ppm (1σ, 1s) | Min. Repeatable Detectable Level: ≤ 0.4 ≤ 0.15 with 10s averaging
Linearity error | <±1%F.S. | < 0.15% F.S + 0.35% for CH4; < 0.4% F.S + 0.1% for CO2 |  | <2% of measuring range. |  |  | < +/- 1% FS
Accuracy | <0.5% F.S. at zero point; Cross-sensitivity: < 2% F.S., to each calibrated gas channel, other gasses on request. | < 1% F.S for CH4 < 2% F.S for CO2 Detection limit: < 2 ppm for CH4 < 35ppm for CO2 Cross-sensitivity at FS: < 35 ppm for CH4 < 100ppm for CO2 | Not defined under 10°c | Zero point drift: <2% of measusuring range per 24h background measurement interval; No sensitivity drift | Capable of reading as low as 50 ppm, with an accuracy of 5 ppm. Above 1500 ppm, an accuracy of 20 ppm can be achieved, up to at least a maximum of 3000 ppm. | 2-5% | 0.5% FS |  | Resolution: 0.01 for CH4 and 5 ppm +2% reading for CO2
Dimensions in cm (WxHxD) | 50x30x20 | 48,6x35x15 | 60x40x20 | 45x28,7x16,6 | 15x30 | 48,4x13,9x44 | 31x22x10 | 56x35,5x19 | 26x36x19
Weight in kg | 10 | 10.9 | 15 | 9.4 | 5 | 13 | 18 | 8 | 14
Network access | Ethernet, still checking for wifi possibilities | HTTP(S) GUI over ethernet WiFi with connection box | Build: LAN, connectors for RJ45 (LAN); Gascard NG: onboard true RS232 communications, optional TCP/IP communications protocol and optional on-board LAN support. | Ethernet, WiFi access point anf WiFi station | Ethernet, WiFi | Ethernet | Wifi/internet | Built in wifi card, Ethernet (RJ-45) | FTP Upload: cellular, ethernet or wifi
Other interfaces/connections/options | 3 RVS PG9 wartels; 6 air connections | RS485 for expansions; I/O for external modules (MOO-HERE and MOO-MXP); Multiplexer for dual nose sampling available | 4G router for outside connecting; PLC use the connecting to synchronize with a timeserver to make sure data is stored with correct time stamp. | 6 mm quick connect; USB, Bluetooth, Rremote operable. | Send to Arcoflex cloud server, can be shared or adjusted to any server | USB and optionally Serial over USB, current message or voltage message. Supports MODBUS and AK-protocol. |  | Serial(RS-232) | Designed to hold a dongle, you can use the provider that works best at that point; USB connectivity underneath device
Error notification | Lights on device | Automatic internal diagnostic and reporting through e-mail notifications and alerts logged in the web interface; | Under (below 4 mA) and over (above 20mA) range gives fault reading. If pump is blocked a fault relay will activate. |  | Send along with data. |  |  |  | 2 indicator lights at the bottom of the device. 1 is for cellular connection.
Calibration | Daily zero calibrations. Span calibrations 6-12 months. | Guided Manual calibration procedure via web interface; Automatic recalibration available | 1x per year CO2. 1x/2-3 months for CH4 (estimated) | Daily manual calibrations with nitrogen. |  |  |  | Yearly | Annual calibration is adviced
Internal data storage | Yes, SD kaart included in Siemens LOGO! | Yes, internal M2 storage; Storage up to 1 year | Denmark: A PLC stored the measurement and handle the multiplexer etc. The measurement is stored in CSV-files; Datalogger with a 176GB SD card |  |  |  |  | 30GB | Optional: 3 years of logs onboard with unique file name creation every 24 hours
Display | Yes, CH4 and CO2 concentrations shown | No | No display visible on outside of build. Only on Guardian inside and only CH4. (29) |  | No | 7 inch WSVGA |  |  | Yes, CH4, CO2, Temperature, humidity and pressure
Temperature range (°C) | 5 °C – 45 °C, for the sensor itself; gas is cooled when going in, build is heated till 40°C; storage between: -20 °C – 60 °C | 0°C - 45°C | 0°C- 40°C (Extensive temperature compensation.) | short term -5°C - 40°C. long term 5°C - 30°C | Used in New Zealand and Australia without problem. | -20°C - 60°C | -20°C to 50 ℃ (-4 to 120 ℉) | 0°C - 45°C | -10°C - 50°C
Temperature stability | Housing heated till 40 °C; Temp influence Zero: < 1% F.S./10K; temp influences span: < 1% F.S./10K, with span temp compensation. | Internal heated and thermalized sensors, stable within operating temperature | If the outside temperature rises above 40° the effect of the sensor heater is void, and the IR-bench will return a fault signal but will not shut down. Below 10° but above freezing will have no effect, except the precision and repeatability is not defined so the uncertaininty of the reading is compromised. Below freezing there’s a risk that pipes will clog from ice crystals due to condensation – and that will return a pump error, but again, the system will not shut down. | Temperature drift < 1 % of measuring range per 10 K temperature change.* Ambient temperature changes are measured and compensated. (* = Typical GHG Application.) |  | Stable within operating temperature.
Humidity | 0 – 95% rel. humidity (not condensing) for the sensor. De peltierkoeler zal de lucht koelen tot ongeveer 7-8 graden °C | 0-95% RH, non-condensing internal gas dryer and humidity monitor |  |  |  | <90% RH, non-condensing; Storage: Non-condensing, maximum relative humidity 99% for temperatures up to 30°C, decreasing linearly to 35% relative humidity at 49°C |  | 0-95% RH, non-condensing
Pressure range | 600 – 1200 hPa (mbar) | 300 - 1200 hPa (mBar) pressure compesated measures | Gascard NG: On-board barometric pressure correction in the range 800mbar to 1150mbar. | Ambient pressure |  | Ambient level; Storage: 750-1050 |  |  | 860 - 1100 hPa
Pressure stability | < 1.5%/10hPa of reading, without pressure compensation | < 1 ppm/10hPa for CH4 < 10ppm/10hPa for CO2 |  | 1 % change of measuring value for 1 % sample pressure change. Ambient pressure changes are measured and compensated. |  | sample gas pressure change within the pressure range will not cause drift
Dust/water resistance | IP53 | IP65 | Depends on build, Bramtech and Duotech: IP65 | IP54 | Cleaning with high pressure hose should have no adverse effect | IP20 (IEC 529) |  |  | IP67
Installation | On top of robot or against wall | Near milking robot, automatic feeders or adapted feeding bunks | On top/next to robot or wall mount; Access door opens to the side. | Has to be used standing up, using it laying down is not recommended | Bolted/stuck to robot, above nose of animal. | Table top and rack mounting | Wall/pole mounting |  | Wall-mount: +- 70 mm offset requiered for dust inlet clearance
Blow back system to clean tube | Optional | Yes, requieres external compressed air and MOO-HERE sensor; | Yes, the gas sniffer is mounted with sensors to measure when the exit door is opening, this information is stored and used for making a backflow in the suction hose. Bramtech build |  | No, no tubes used |  |  |  | No
Own (web) application needed | Siemens LOGO! Interface | Yes, WEB GUID for management, data export, calibration. API for third part systems integrations | It is possible to connect the gas sniffer by Team Viewer | Yes, Calcmet | Yes |  | Internet based user interface |  | Yes
Cloud service |  | Optional | The datafiles are send by FTP / OneDrive cloud and stored on SQL – server |  | Yes |  | Yes
RFID (automatically assign data to individual animals) | No | Optional | No | No | Optional | No | Yes, reads only half duplex and FDX-full duplex tags | No | Optional
Head position sensor | No | No | No | No | No | No | Yes | No | Based on distance of tag the reader
Condensation | Heat sink | Water detection system with water block, discharge and automatic dry for autonomous restart | A Perma Pure MD-110-12F-4 to dry the air from the feed bin.<- From Denmark. From DuoTec-> The sensor heater in the IR-bench operates at 45° to ensure no condensation is formed on the detector filter.
Power consumption | On average <2W; 24V | Typ. 20-75W (depending on environmental conditions); Max.150 W | gascard NG: Minimum operating voltage 7V and wide operating voltage range (7V to 30V). | Li-ion battery, approx. 3 hours of operation |  |  | Can operate from 12 volt DC, 120 volt AC, or 240 AC power | < 30 W | 43W
Power requirements/supply | 230V; 24 (15 – 30) VDC; suply current (peak): <0,1 A | 100-120 VAC, 60Hz or 220-240 VAC, 50 Hz | 230 VAC | 115 / 230 VAC |  | input voltage: 90…264 Vac, 47…63 Hz Input power: 75 W max | 100 ~ 240VAC / 2.8A 10.5 ~ 15VDC / 10A | 14 - 28 VDC, 2,5 A (33) | 12VDC
Maintenance by company | Initially 6 months maintenance (estimate) | Optional | No | [5 empty] | Laser Module Operating Life 8000 hours
Delivery time | 1 month | 1-3 months | 8-10 weeks | 5-7 weeks | 8 weeks |  |  |  | 8 weeks
Price range | low | low | low | high | low | high | high | high | mid